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D260BLX Brushless Micropump Range
TCS MICROPUMPS: D260BLX (RANGE)
(Patents Pending, Quality Assured ISO 9001, RoHS compliant, IP41 protection rating)


INTRODUCTION
The TCS D series Micropumps are high-quality self priming miniature pumps for gases. They are highly efficient, small and lightweight. The solid construction and wide temperature tolerance, enable them to perform reliably even in hostile environments. Your Micropump can be quickly and easily installed into the smallest spaces, in a vast range of laboratory, prototype and production equipment.

ELECTRICAL CONNECTION
	
CAUTION: REVERSED OR INCORRECT CONNECTION WILL PERMANENTLY DAMAGE THE ELECTRONICS IN YOUR PUMP!

Voltage: 3.0 to 6.0 Volts dc (Maximum 0.5Amps)

The simplest connection is Red, Green and Grey together to +ve (Positive) supply, Black to –ve (Negative) of supply


	Wire Colour
	Description
	Connection
	Status

	Black
	GRD
The negative connection
	-Ve
	Required

	Grey
	VM 
Motor Power Supply
	+Ve   3.0~6.0 
	Required

	Red
	Vcc
Onboard Controller Supply Circuit
	+Ve   3.0~6.0
	Required

	Green
	CW/CCW
Motor Rotation
	+Ve   3.0~6.0
	Required

	Yellow
	RPM
Signal Wire
	GND
Vcc 4 pulse / revolution 
square wave signal
	Optional



Your Micropump can be connected to batteries, a DC power source of up to 6V or a TCS power supply TCS-600 (see accessories page on website).

Pumping rate can be adjusted by varying input voltages between 3.0v and 6.0v dc. Further fine tuning can be made by restricting the outlet of your pump using a fixed jet restriction or flexible tubing and an adjustable clamp. The pump is quite happy to run even with the outlet completely blocked with gases.

Materials
	Housing:	Stainless Steel 316 (D260) 
Tubing Connectors:	Stainless Steel 316 (D260) 
Seals:	(Viton + Viton), (EPDM + EPDM), (Silicone + Nitrile), (Silicone + Viton)


TUBING
D200 series:	2.4mm (3/32”) bore flexible tubing 


Notes on Operation
All ‘D series’ pumps are self priming by design. Simply connect the pump into your flow circuit, noting the flow direction shown below, and begin pumping.
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                                              D260BLX



	Performance
	Input Voltage
	Power Usage
	Max Pressure
	Max Free Flow

	
	(v)
	(w)
	(psi)
	(mlpm)

	D260BLX - AIR
	3.0 - 6.0
	0.1 - 0.6
	11
	550



Please note that figures above may vary depending on elastomers featured in pump.
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